Low battery warning:

- If the battery load is < 2.70 V, the low battery sign blinks.

- If the battery load is < 2.60 V, the outer frame of battery symbol is displayed with . With that, no backlight is possible and no measurement can be taken. The device turns off automatically after 30s.

End of measurement:

Object mode – The end of a measurement is signalized through sound/vibration (“-“). The backlight is always green.

Forehead mode/ ear mode – The end of a measurement is signalized through sound/vibration (see table below).

Together with the (alarm) signal/vibration, a colored backlight turns on (green, orange, red). After 15 s, the colored backlight turns off again. For auditive/haptic feedback, the following beep/vibration pattern are implemented:

| Temperature [°C] |  | Temperature [°F] | Beep/Vibration pattern | Backlight color | Smiley indicator | Display |
| --- | --- | --- | --- | --- | --- | --- |
|  | 0 – 3 months |  |  |  |  |  |
| 34.0 ≤ T ≤ 37.4 |  | 93.2 ≤ T ≤ 99.4 | - | Green |  |  |
| 37.4 < T ≤ 43.0 |  | 99.4 < T ≤ 109.4 | - · · · | Red |  |  |
| T < 34.0 |  | T < 93.2 | · · · · | Green |  | Lo |
| T > 43.0 |  | T > 109.4 | - · · · | Red |  | Hi |
|  | 3 – 36 months |  |  |  |  |  |
| 34.0 ≤ T ≤ 37.6 |  | 93.2 ≤ T ≤ 99.6 | - | Green |  |  |
| 37.6 < T ≤ 38.5 |  | 99.6 < T ≤ 101.3 | - | Orange |  |  |
| 38.5 < T ≤ 43.0 |  | 101.3 < T ≤ 109.4 | - · · · | Red |  |  |
| T < 34.0 |  | T < 93.2 | · · · · | Green |  | Lo |
| T > 43.0 |  | T > 109.4 | - · · · | Red |  | Hi |
|  | 36+ months |  |  |  |  |  |
| 34.0 ≤ T ≤ 37.7 |  | 93.2 ≤ T ≤ 99.9 | - | Green |  |  |
| 37.7 < T ≤ 39.4 |  | 99.9 < T ≤ 103.0 | - | Orange |  |  |
| 39.4 < T ≤ 43.0 |  | 103.0 < T ≤ 109.4 | - · · · | Red |  |  |
| T < 34.0 |  | T < 93.2 | · · · · | Green |  | Lo |
| T > 43.0 |  | T > 109.4 | - · · · | Red |  | Hi |

Note：“-” long beep (250ms)，“·” short beep (180ms)

Automatic switch-off: After 30 s with no button press, the device turns off automatically (without any vibration/sound feedback).

The key flow described below applies in every "age mode."

| Button (left = age/MEM, right = settings, middle = On/OFF) | Button press duration (short: <1s, long: ca. 3s) | Device mode | Sound mode | Measurement mode | Cap mode | Description |
| --- | --- | --- | --- | --- | --- | --- |
| ☒left button ☐middle button ☐right button | ☒short ☐long | ☐on ☒off | ☒on ☒off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☐cap on ☐cap off | Memory mode |
| ☐left button ☒middle button ☐right button | ☐short ☒long | ☐on ☐off ☒on, memory mode | ☐on ☒off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☐cap on ☐cap off | Turn device off by long button press for 4s. Vibration feedback (“-“) signalizes successful turn-off. |
| ☐left button ☒middle button ☐right button | ☐short ☒long | ☐on ☐off ☒on, memory mode | ☒on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☐cap on ☐cap off | Turn device off by long button press for 4s. Sound feedback (“-“) signalizes successful turn-off. |
| ☒left button ☐middle button ☐right button | ☒short ☐long | ☒on ☐off | ☐on ☒off (Vibration) | ☒Forehead mode ☐Object mode ☒Ear mode | ☒cap on ☒cap off | Device must vibrate if age function mode is changed by button press. |
| ☒left button ☐middle button ☐right button | ☒short ☐long | ☒on ☐off | ☒on ☐off (Vibration) | ☒Forehead mode ☐Object mode ☒Ear mode | ☒cap on ☒cap off | Device must give signal if age function mode is changed by button press. |
| ☒left button ☐middle button ☐right button | ☒short ☐long | ☒on ☐off | ☒on ☒off (Vibration) | ☒Forehead mode ☐Object mode ☒Ear mode | ☒cap on ☒cap off | Switching-time between age modes should be the same as in sound-on mode, to ensure consistency between sound-on/sound-off vibration modes. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☐on ☒off | ☐on ☒off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☐cap on ☐cap off | Turn device on without vibration feedback and when ready to measure “double vibration” as it is in sound mode but here with vibration. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☐on ☒off | ☒on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☐cap on ☐cap off | Turn device on without sound feedback and when ready to measure “double beep”. |
| ☐left button ☒middle button ☐right button | ☐short ☒long | ☒on ☐off | ☐on ☒off (Vibration) | ☒Forehead mode ☒Object mode ☒Ear mode | ☒cap on ☒cap off | Turn device off by long button press for 4s. Vibration feedback (“-“) signalizes successful turn-off. |
| ☐left button ☒middle button ☐right button | ☐short ☒long | ☒on ☐off | ☒on ☐off (Vibration) | ☒Forehead mode ☒Object mode ☒Ear mode | ☒cap on ☒cap off | Turn device off by long button press for 4s. Sound feedback (“-“) signalizes successful turn-off. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☒on ☐off | ☒on ☐off (Vibration) | ☒Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☐cap off | Swipe-measurement starts (3s), rotating icon on display shows ongoing measurement, sound feedback + 15s backlight when measurement finished, show highest value. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☒on ☐off | ☐on ☒off (Vibration) | ☒Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☐cap off | Swipe-measurment starts (3s), rotating icon on display shows ongoing measurement, vibration feedback + 15 s backlight when measurement finished, show highest value. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☒on ☐off | ☒on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☒Ear mode | ☐cap on ☒cap off | Start measurement, sound feedback + 15 s backlight when measurement finished, after 5 seconds “ready-to-measure” double-beep feedback1. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☒on ☐off | ☐on ☒off (Vibration) | ☐Forehead mode ☐Object mode ☒Ear mode | ☐cap on ☒cap off | Start measurement, vibration feedback + 15 s backlight when measurement finished. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☒on ☐off | ☒on ☐off (Vibration) | ☐Forehead mode ☒Object mode ☐Ear mode | ☒cap on ☐cap off | Start measurement (1s), sound feedback when measurement finished. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☒on ☐off | ☐on ☒off (Vibration) | ☐Forehead mode ☒Object mode ☐Ear mode | ☒cap on ☐cap off | Start measurement (1s), vibration feedback when measurement finished. |
| ☐left button ☐middle button ☒right button | ☒short ☐long | ☒on ☐off | ☐on ☐off (Vibration) | ☒Forehead mode ☒Object mode ☒Ear mode | ☒cap on ☒cap off | Sound on/off, signalized by sound or vibration + icon in display if sound on. |
| ☐left button ☐middle button ☒right button | ☐short ☒long | ☒on ☐off | ☒on ☒off (Vibration) | ☒Forehead mode ☒Object mode ☐Ear mode | ☒cap on ☐cap off | Switch between forehead and object mode. |
| ☐left button ☐middle button ☒right button | ☐short ☒long | ☒on ☐off | ☒on ☒off (Vibration) | ☐Forehead mode ☐Object mode ☒Ear mode | ☐cap on ☒cap off | For ear mode, no action with this kind of button press. |
| ☐left button ☐middle button ☒right button | ☐short ☒long | ☐on ☒off | ☐on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☒cap off | Show °C/°F menu. |
| ☐left button ☒middle button ☐right button | ☒short ☐long | ☐on ☐off ☒off °C/°F-mode | ☐on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☒cap off | Change between °C/°F. |
| ☐left button ☐middle button ☒right button | ☒short ☐long | ☐on ☐off ☒ off °C/°F-mode | ☐on ☒off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☒cap off | Select °C/°F, turn device off (with vibration feedback “-“). |
| ☐left button ☐middle button ☒right button | ☒short ☐long | ☐on ☐off ☒ off °C/°F-mode | ☒on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☒cap off | Select °C/°F, turn device off (with sound feedback “-“). |
| ☐left button ☐middle button ☐right button | ☐short ☐long | ☒on ☐off | ☐on ☒off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☒cap off | When cap put on/off the device, no vibration feedback. |
| ☐left button ☐middle button ☐right button | ☐short ☐long | ☒on ☐off | ☒on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☒cap on ☒cap off | When cap put on/off the device, no sound feedback. |
| ☐left button ☒middle button ☐right button | ☐short ☒long (7s) | ☒on ☐off | ☐on ☐off (Vibration) | ☐Forehead mode ☐Object mode ☐Ear mode | ☐cap on ☐cap off | Access calibration mode, “.” appears as soon as calibration mode is entered |

| RELEASE |  |
| --- | --- |
| Issued / changed by | Checked / approved by |

| HISTORY | Changes to the previous document will be marked in red and documented below. Before updating the document, all previous red markings have to be set back to black colour. |  |  |
| --- | --- | --- | --- |
| Revision | Date | Description | Issued / changed by |
| 01 | 2024-10-24 | Initial version | Antonia Bernecker |
| 02 | 2024-11-26 | Added low battery information; updated end of measurement specification; updated Keyflow according to development decisions | Antonia Bernecker |
| 03 | 2025-xx-xx | Updated prolonged sound/vibration duration due to vibration motor specification; added loading-symbol information during forehead measurement; added information on entering calibration mode | Antonia Bernecker |